## Sheet1
Function | Hazard | Cause | SafetyMessage | HazardKey
outside influences / environment | heart arrhythmia | summation of patient leakage current prevents correct function of additional external ECG equipment | Operator Manual: WARNING Source: Exceeded safety limits of total leakage current when using different electronic devices on one patient in combination with the CT scanner! Consequences: Heart arrhythmia Countermeasure: Do not use more electronic devices on the patient than absolutely necessary. Always consider the total leakage current when using external electronic devices in combination with the CT scanner. The correct functionality of your external equipment may also be affected, for example, ECG equipment. | {hz_ga_F14G09U01M01}
service hazard | Damage to the CT system | Tablet not in the cradle during update | Operator Manual: CAUTION Source: Tablet not in the cradle during updates! Consequence: Wrong installation and damage to the equipment. Countermeasure: Only download and install a BIOS update when the tablet is in the cradle. Do not shut down or restart the CT system. | {hz_ga_F13G13U05M01}
outside influences / environment | heart arrhythmia | summation of patient leakage current prevents correct function of additional external ECG equipment | Operator Manual: WARNING Source: Exceeded safety limits of total leakage current when using different electronic devices on one patient in combination with the CT scanner! Consequences: Heart arrhythmia Countermeasure: Do not use more electronic devices on the patient than absolutely necessary. Always consider the total leakage current when using external electronic devices in combination with the CT scanner. The correct functionality of your external equipment may also be affected, for example, ECG equipment. | {hz_ga_F14G09U01M01}
outside influences / environment | damage to people and property, injury by gaspoisoning due to burning plastic | fire arround the system | Operator Manual: WARNING Source: Fire inside or in the vicinity of the system! Consequences: Injury to the personnel or damage to the equipment. Gas poisoning due to burning plastic. Countermeasure: Switch off the system in the event of fire. Make sure that you know where the escape routes are. Make sure that you know where the fire extinguishers are located and familiarize yourself with the use of them. | {hz_ga_F14G03U01M05}
archiving | Data lost | Low-quality storage media are used. | Operator Manual: CAUTION Source: Low-quality storage media are used! Consequence: Data loss due to improper handling or low-quality CDs or DVDs. Countermeasure: Use only quality media approved for medical applications when storing patient data. Observe proper handling guidelines for CDs and DVDs. | {hz_ga_F08G02U06M01}
Scanner Movement and positioning | (ICU) System is driven on a ramp | Damage to the system or serious injury | Operator Manual: WARNING Source: The system is driven on a ramp without proper preparation! Consequence: Damage to the system or serious injury. Countermeasure: Before driving the system, ensure it's properly prepared and follow the Instructions for Use. Check that the gantry is centered, do not use the gantry bore to transport objects/persons, and only move the system while the motor drive is engaged. | {hz_ga_F25G01U14M02}
patient registration (POF) | incorrect diagnosis basis | wrong entry of patient orientation | Operator Manual: CAUTION Source: Wrong entry of patient position! Consequence: Wrong basis for diagnosis. Countermeasure: Make sure that the patient position is correct. | {hz_ga_F01G01U03M03}
archiving | Data lost | Low-quality storage media are used. | Operator Manual: CAUTION Source: Low-quality storage media are used! Consequence: Data loss due to improper handling or low-quality CDs or DVDs. Countermeasure: Use only quality media approved for medical applications when storing patient data. Observe proper handling guidelines for CDs and DVDs. | {hz_ga_F08G02U06M01}
network & communication | Images seem to be misaligned or lost | Changes of patient information on other systems will not automatically change the patient information in syngo! | Operator Manual: CAUTION Source: If syngo receives images but the study or series UID is already available in the local database, the data will be assigned to this locally existing patient. This happens regardless of the patient identification (name, date of birth, gender and patient ID) of the received images. Changes of patient information on other systems will not automatically change the patient information in syngo! Consequence: Data retrieved from PACS or other nodes seems to be misaligned or lost, but is stored elsewhere in the database. Countermeasure: Imported data will always be appended to the already existing patient on the local system, based on matching DICOM UIDs. If you imported data, but cannot find them on the local system under a certain name, use other criteria from the study or series to search for the imported data, or filter the data. To completely avoid these inconsistencies, always delete the local patient data after you have successfully archived them. | {hz_ga_F11G18U01M01}